# Supplementary

Figure S1: Axial and sagittal views through the center of the heart showing B1+ efficiency maps (B1+ /√1kW) obtained for an equal phase (0) and equal amplitude (1) excitation using the a) self-grounded bow-tie antenna building block; b) bow-tie antenna building block; and c) the fractionated dipole antenna RF array configurations at 7.0 T (baseline, BL) and 14.0 T (same channel count, SCC, double channel count, DCC). The cardiac ROI is depicted in red. The spider diagrams illustrate the relative changes of the mean B1+ROI, minimum B1+ROI, maximum SAR10g, CoV(B1+ROI), TXE, and intrinsic SNR values for the 14.0 T SCC (orange) and DCC (grey) with respect to the 7.0 T baseline(black).

Table S1 Summary of the mean B1+, minimum B1+ , coefficient of variation (CoV(B1+ROI)) across the entire heart of the human voxel models Duke and Ella, and the maximum SAR10g for an excitation vector with equal phase (0°) and amplitude (1) for all channels using self-grounded bow-tie (SGBT) antenna building block, bow-tie (BT) antenna building block, and fractionated dipole (FD) antenna RF arrays at 7.0 T (baseline, BL) and 14.0 T (same channel count, SCC, double channel count, DCC). The total RF power for the excitation vectors (Pfwd) is presented for a lossless 2kW power at each channel.

| Excitation with equal phase (0°) and amplitude (1) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Duke |  | mean B1+ROI [µT/√kW] | min. B1+ROI [µT/√kW] | max. SAR10g [W/kg] | CoV [%] | Pfwd [kW] |
| 7.0 T BL | SGBT | 5.21 | 0.02 | 0.30 | 45 | 64 |
|  | BT | 3.40 | 0.03 | 0.23 | 44 | 32 |
|  | FD | 5.05 | 0.01 | 0.25 | 39 | 16 |
| 14.0 T SCC | SGBT | 3.37 | 0.05 | 0.62 | 52 | 64 |
|  | BT | 1.96 | 0.01 | 0.67 | 55 | 32 |
|  | FD | 3.38 | 0.04 | 0.57 | 48 | 16 |
| 14.0 T DCC | SGBT | 4.58 | 0.03 | 0.45 | 56 | 128 |
|  | BT | 2.80 | 0.03 | 0.25 | 46 | 64 |
|  | FD | 2.51 | 0.03 | 0.41 | 48 | 32 |
| Ella |  |  |  |  |  |  |
| 7.0 T BL | SGBT | 5.41 | 0.06 | 0.29 | 39 | 64 |
|  | BT | 4.31 | 0.04 | 0.19 | 41 | 32 |
|  | FD | 6.17 | 0.06 | 0.28 | 39 | 16 |
| 14.0 T SCC | SGBT | 3.70 | 0.01 | 1.08 | 51 | 64 |
|  | BT | 2.44 | 0.02 | 0.50 | 44 | 32 |
|  | FD | 3.51 | 0.02 | 0.40 | 41 | 16 |
| 14.0 T DCC | SGBT | 4.85 | 0.03 | 0.49 | 52 | 128 |
|  | BT | 3.21 | 0.04 | 0.22 | 39 | 64 |
|  | FD | 2.79 | 0.03 | 0.27 | 42 | 32 |

Figure S2: Axial and sagittal views through the center of the heart showing B1+ efficiency maps (B1+ /√1kW) obtained for static pTx phase shimming using the a) self-grounded bow-tie antenna building block; b) bow-tie antenna building block; and c) the fractionated dipole antenna RF array configurations at 7.0 T (baseline, BL) and 14.0 T (same channel count, SCC, double channel count, DCC). The cardiac ROI is depicted in red. The superposed minimum B1+ of all channels within the whole 3D cardiac ROI was maximized in the optimization process. The spider diagrams illustrate the relative changes of the mean B1+ROI, minimum B1+ROI, maximum SAR10g, CoV(B1+ROI), TXE, and intrinsic SNR values for the 14.0 T SCC (orange) and DCC (grey) with respect to the 7.0 T baseline(black).

Table S2: Summary of the mean B1+, minimum B1+, coefficient of variation (CoV(B1+ROI)) across the entire heart of the human voxel models Duke and Ella, and the maximum SAR10g for a phase pTx approach with optimized minimum B1+ in the ROI self-grounded bow-tie (SGBT) antenna building block, bow-tie (BT) antenna building block, and fractionated dipole (FD) antenna RF arrays at 7.0 T (baseline, BL) and 14.0 T (same channel count, SCC, double channel count, DCC). The total RF power for the excitation vectors (Pfwd) is presented for a lossless 2kW power at each channel.

| Minimum B1+ Optimization |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Duke |  | mean B1+ROI [µT/√kW] | min. B1+ROI [µT/√kW] | max. SAR10g [W/kg] | CoV [%] | Pfwd [kW] |
| 7.0 T BL | SGBT | 5.85 | 3.16 | 0.35 | 34 | 64 |
|  | BT | 2.87 | 1.46 | 0.31 | 34 | 32 |
|  | FD | 5.49 | 2.57 | 0.41 | 32 | 16 |
| 14.0 T SCC | SGBT | 3.06 | 0.10 | 0.72 | 59 | 64 |
|  | BT | 1.82 | 0.34 | 0.81 | 60 | 32 |
|  | FD | 3.53 | 0.22 | 0.76 | 51 | 16 |
| 14.0 T DCC | SGBT | 2.66 | 0.45 | 0.86 | 70 | 128 |
|  | BT | 2.68 | 0.64 | 0.44 | 51 | 64 |
|  | FD | 3.39 | 0.39 | 0.57 | 65 | 32 |
| Ella |  |  |  |  |  |  |
| 7.0 T BL | SGBT | 6.69 | 4.17 | 0.35 | 24 | 64 |
|  | BT | 3.71 | 1.93 | 0.34 | 27 | 32 |
|  | FD | 6.99 | 3.71 | 0.36 | 25 | 16 |
| 14.0 T SCC | SGBT | 4.01 | 1.12 | 0.81 | 54 | 64 |
|  | BT | 2.12 | 0.58 | 0.50 | 42 | 32 |
|  | FD | 4.51 | 0.52 | 0.60 | 55 | 16 |
| 14.0 T DCC | SGBT | 4.65 | 1.33 | 0.77 | 52 | 128 |
|  | BT | 2.87 | 0.89 | 0.54 | 37 | 64 |
|  | FD | 3.47 | 0.69 | 0.72 | 50 | 32 |

Figure S3: Axial and sagittal views through the center of the heart showing B1+ efficiency maps (B1+ /√1kW) obtained for static pTx phase shimming using the a) self-grounded bow-tie antenna building block; b) bow-tie antenna building block; and c) the fractionated dipole antenna RF array configurations at 7.0 T (baseline, BL) and 14.0 T (same channel count, SCC, double channel count, DCC). The cardiac ROI is depicted in red. The CoV (B1+) within the whole 3D cardiac ROI was minimized in the optimization process. The spider diagrams illustrate the relative changes of the mean B1+ROI, minimum B1+ROI, maximum SAR10g, CoV(B1+ROI), TXE, and intrinsic SNR values for the 14.0 T SCC (orange) and DCC (grey) with respect to the 7.0 T baseline(black). The total RF power for the excitation vectors (Pfwd) is presented for a lossless 2kW power at each channel.

Table S3: Summary of the mean B1+, minimum B1+, coefficient of variation (CoV(B1+ROI)) across the entire 3D heart of the human voxel models Duke and Ella, and the maximum SAR10g for a phase pTx approach with optimized CoV(B1+ROI) in the ROI using self-grounded bow-tie (SGBT) antenna building block, bow-tie (BT) antenna building block, and fractionated dipole (FD) antenna RF arrays at 7.0 T (baseline, BL) and 14.0 T (same channel count, SCC, double channel count, DCC). The total RF power for the excitation vectors (Pfwd) is presented for a lossless 2kW power at each channel.

| CoV Optimization |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Duke |  | mean B1+ROI [µT/√kW] | min. B1+ROI [µT/√kW] | max. SAR10g [W/kg] | CoV [%] | Pfwd [kW] |
| 7.0 T BL | SGBT | 2.99 | 1.06 | 0.56 | 15 | 64 |
|  | BT | 1.93 | 0.75 | 0.37 | 21 | 32 |
|  | FD | 5.29 | 1.70 | 0.34 | 28 | 16 |
| 14.0 T SCC | SGBT | 3.49 | 0.04 | 0.61 | 42 | 64 |
|  | BT | 1.06 | 0.02 | 0.83 | 40 | 32 |
|  | FD | 3.36 | 0.08 | 0.74 | 44 | 16 |
| 14.0 T DCC | SGBT | 2.04 | 0.02 | 0.82 | 33 | 128 |
|  | BT | 1.37 | 0.02 | 0.30 | 32 | 64 |
|  | FD | 2.50 | 0.09 | 0.53 | 40 | 32 |
| Ella |  |  |  |  |  |  |
| 7.0 T BL | SGBT | 4.31 | 2.16 | 0.78 | 15 | 64 |
|  | BT | 2.73 | 0.87 | 0.21 | 19 | 32 |
|  | FD | 6.25 | 2.02 | 0.43 | 21 | 16 |
| 14.0 T SCC | SGBT | 2.77 | 0.06 | 0.95 | 34 | 64 |
|  | BT | 1.41 | 0.02 | 0.49 | 33 | 32 |
|  | FD | 3.69 | 0.08 | 0.64 | 35 | 16 |
| 14.0 T DCC | SGBT | 1.93 | 0.54 | 0.81 | 28 | 128 |
|  | BT | 2.00 | 0.08 | 0.56 | 27 | 64 |
|  | FD | 3.19 | 0.01 | 0.62 | 34 | 32 |

Figure S4: a-c) Pareto front derived from the static phase optimized pTx MOO approach using the a) self-grounded bow-tie antenna building block; b) bow-tie antenna building block; and c) the fractionated dipole antenna RF array configurations at 7.0 T (baseline, BL) and 14.0 T (same channel count, SCC, double channel count, DCC). Each point of the solution represents one optimized excitation vector where a trade-off between the minimum B1+ROI and the maximum SAR10g was found. The green circles indicate the highest minimum B1+ROI/√SAR10g ratio d-f): Axial and sagittal views through the center of the heart (depicted in red) illustrating B1+ efficiency maps (B1+ /√1kW) obtained for the excitation vectors  with highest minimum B1+ROI/√SAR10g ratio (indicated by the green circles in a-c). The spider diagrams illustrate the relative changes of the mean B1+ROI, minimum B1+ROI, maximum SAR10g, CoV(B1+ROI), TXE, and intrinsic SNR values for the 14.0 T SCC (orange) and DCC (grey) with respect to the 7.0 T baseline (black).

Table S4: Summary of the mean B1+, minimum B1+, coefficient of variation (CoV(B1+)) across the entire 3D heart of the human voxel models Duke and Ella, and the maximum SAR10g for the multiobjective phase optimizer with a trade-off between minimum B1+ROI in the heart and maximum SAR10g using self-grounded bow-tie (SGBT) antenna building block, bow-tie (BT) antenna building, and fractionated dipole (FD) antenna RF arrays at 7.0 T (baseline, BL) and 14.0 T (same channel count, SCC, double channel count, DCC). The total RF power for the excitation vectors (Pfwd) is presented for a lossless 2kW power at each channel.

| Multiobjective Optimization |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Duke |  | mean B1+ROI [µT/√kW] | min. B1+ROI [µT/√kW] | max. SAR10g [W/kg] | CoV [%] | Pfwd [kW] |
| 7.0 T BL | SGBT | 5.17 | 1.17 | 0.33 | 44 | 64 |
|  | BT | 2.98 | 1.10 | 0.09 | 40 | 32 |
|  | FD | 5.41 | 2.19 | 0.29 | 34 | 16 |
| 14.0 T SCC | SGBT | 3.39 | 0.03 | 0.56 | 49 | 64 |
|  | BT | 1.94 | 0.05 | 0.28 | 58 | 32 |
|  | FD | 3.68 | 0.11 | 0.61 | 50 | 16 |
| 14.0 T DCC | SGBT | 4.50 | 0.08 | 0.37 | 54 | 128 |
|  | BT | 2.59 | 0.28 | 0.32 | 45 | 64 |
|  | FD | 3.19 | 0.38 | 0.40 | 55 | 32 |
| Ella |  |  |  |  |  |  |
| 7.0 T BL | SGBT | 5.45 | 2.10 | 0.35 | 34 | 64 |
|  | BT | 3.09 | 1.11 | 0.16 | 34 | 32 |
|  | FD | 6.93 | 3.40 | 0.27 | 26 | 16 |
| 14.0 T SCC | SGBT | 3.48 | 0.06 | 0.51 | 48 | 64 |
|  | BT | 2.31 | 0.36 | 0.34 | 41 | 32 |
|  | FD | 3.83 | 0.39 | 0.52 | 53 | 16 |
| 14.0 T DCC | SGBT | 4.74 | 0.04 | 0.43 | 53 | 128 |
|  | BT | 3.14 | 0.04 | 0.19 | 41 | 64 |
|  | FD | 3.06 | 0.11 | 0.27 | 49 | 32 |